## Worksheet
Digital Pathology QMSR Release-Readiness Checklist
Editable preparation aid — not an FDA form or certification tool. Complete and review against applicable requirements.
Checklist area | Evidence to prepare | Owner | Gate | Status | Release test | Linked records | Risk note
Intended use and classification | Document users, specimens, inputs, outputs, adjunctive role, classification and authorization basis | Regulatory lead | Design input | Not started | Challenge the claim against the controlled intended-use statement | Regulatory assessment; labeling; requirements | An overbroad claim changes evidence needs
Compatibility matrix | Control scanners, viewing software, specimen types, operating environments and user roles | Product lead | Before verification planning | Not started | Select one configuration and trace every compatibility claim | Requirements; labeling; test plan | Untested combinations can invalidate workflow claims
Software lifecycle | Connect requirements, architecture, risk controls, verification, validation, release and feedback | Software quality lead | Each release | Not started | Trace one requirement through release | DHF; risk file; release record | Disconnected artifacts weaken design-control evidence
Technical performance | Separate image acquisition, viewing, processing and system-performance evidence from clinical evidence | Verification lead | Before clinical review | Not started | Review whether each technical claim has objective evidence | Protocol; report; anomaly log | Technical success does not establish clinical performance
Clinical evaluation | Document valid clinical association, analytical validation and clinical validation | Clinical lead | Before submission or claim change | Not started | Map every clinical claim to its supporting evidence | Clinical evaluation plan; study report | Mixed evidence categories obscure gaps
Cybersecurity applicability | Record software presence, connectivity capability, vulnerability characteristics and section 524B determination | Security lead | Design input and each architecture change | Not started | Reassess one network-connected configuration | Threat model; applicability memo | Connectivity changes can alter obligations
SBOM and vulnerability response | Maintain component inventory, monitoring ownership, triage criteria, remediation and communication | Security lead | Continuous | Not started | Trace one known component through monitoring and disposition | SBOM; vulnerability log; patch plan | Unknown components delay response
Change control | Assess requirements, risk, performance, cybersecurity, labeling, V&V and submission impact before release | Change board | Every software change | Not started | Run one historical change through the current gate | Change record; regression evidence; regulatory memo | Undocumented decisions are difficult to defend
QMS software assurance | Distinguish assurance for production and quality-system software from validation of the marketed device software | IT quality lead | Tool onboarding and change | Not started | Review one eQMS or automated quality tool | Intended use; risk assessment; assurance record | Wrong scope creates either excess work or missing evidence
Cahir MedTech · © Cahir AI LLC 2026